# HOBBY EBERLY TELESCOPE TRACKER

# Y ROLLER-SCREW RETAINER NUT REPLACEMENT PROCEDURE

# April 24, 2002

Revision 1

# Originator

1.0 Purpose

The purpose of this procedure is to provide the information required to replace the Tracker Y roller-screw retainer nut.

# 2.0 Scope

This procedure defines the steps, parts, hardware, and equipment for removing the existing Tracker Y roller-screw retainer nut and installing the new Y roller-screw retainer nut.

# 3.0 Applicable Documents

The following documents are applicable to the extent specified herein:

a. 1010-0020 	Tracker Beam Assembly

b. TM-12774 	HET Project Safety plan

c. TM-12725   HET Tracker Installation Procedure

d. TM-13058	HET Tracker Servicing Procedure

# 4.0 Safety Precautions

Y roller-screw retainer nut replacement will be performed at the University of Texas, McDonald Observatory.  Measures shall be taken to prevent, or reduce to a minimum, personnel exposure to hazardous conditions and environments.  A safe servicing area will be maintained during the Y roller-screw retaining nut replacement.  The following personnel and facility safety precautions shall be observed:

- One engineer shall be designated as the Installation Director and will be responsible for the safety of the personnel and equipment during conduct of the replacement procedure.

- Secure the area from unnecessary traffic.

- Protective equipment shall be worn as required by established safety practice.

- Prior to working with electrical circuits, remove all jewelry, watches, rings, and metal objects.

- Personnel shall read and understand safety precautions outlined in reference documents.

- Personnel involved with elevated work shall wear fall restraint equipment as specified in Section 9 of TM-12774 or per HET facility procedures.

- All personnel on the ground level shall wear protective equipment as defined in TM-12774 per HET facility procedures.

- The Installation Director shall review the complete procedure with the participants.

# 5.0 Required Special Equipment

- One (1) 5000 lb minimum sling. Length to reach the top of the carriage, when at 3 feet from the top of travel, and routed under the top cross member of the Y beam to the dome crane hook.

- Two (2) 3000 lb minimum slings.  3 to 5 feet long.

- Compressed air supply to the reach the Y axis tracking brake.  With fitting to attach to the brake air input.

- Brake caliper shim (Provided with Tracker)

- Three (3) D size sheets of mylar (0.004 to 0.010 thick)

- One (1) wood V-block (Provided with Tracker) or high enough to reach the roller screw at the floating bearing mount (-Y end).  V cut to fit roller screw.

- Bearing retainer installation/removal tool – spanner (Delivered with Tracker)

- Large ratchet driver for item 7.

- Large adjustable wrench and “cheater bar”  (Pipe to extend lever arm of adjustable wrench).

- Torque wrench with a socket driver head (Used on bearing retainer installation/removal tool).  Torque required is minimum 116 Ft-lbs.

# 6.0 Y Roller-Screw Retainer Nut Replacement

## 6.1 Preparation

A minimum of two (2) people is required to perform these operations.  Three (3) people are preferred

- The Installation Director shall review this portion of the procedure with all participants.  All necessary drawings, procedures, equipment, etc shall be available for use during this briefing and the work.

- Load all special equipment required into the manlift.

- Run the compressed air line to the Y track brake.

- Drive the Carriage in the +Y direction to the roller screw fixed bearing mount.

(Track motor).  Apply all brakes.

- Position the dome crane at the top (+Y) end of the Y beam.  Lower the crane hook to within reach of the top of the Y beam.

- Two (2) slings (3000 lb min rating) are required.  Attach one (1) sling at the bottom on each side of the carriage around the Y beam vertical truss member.  Leaving approximately 24 inches of slack to lower the carriage in the –Y direction.

- Attach one (1) sling (5000 lb min rating) around the upper crossmember of the carriage in the center.  Run the sling under the top cross-member X beam of the Y beam and attach that end to the crane.

CAUTION:

THE CARRIAGE MAY BACK-DRIVE DURING THE NEXT STEP.  HAVE ONE PERSON MANING THE BRAKE SET VALVE TO QUICKLY SET THE BRAKE IF NECESSARY.

- Gently pull the strap tight with the crane.

- Release the Y track brake as follows: disconnect the air line to the brake if not already removed; connect the air line from the portable tank to the brake; open the air supply valve until the brake releases then shut the air valve again.  If the carriage begins to back-drive excessively set the brake by opening the vent valve.  Repeat steps 8 and 9 until the crane has the downhill weight of the carriage.

## 6.2 Y Roller-Screw Lower Limit Switch Removal

## At the lower end of the Y Roller-Screw, remove the cover for the Emergency

Stop switch.

2) Inspect the lower end for tool clearance.  Identify the size wrench required for counter-torque.

## 6.3 Y Roller-Screw Retainer Nut Removal

A minimum of two (2) people is required to perform these operations.  Three (3) are preferred.

- Locate the brake caliper spacer provided with the Tracker and keep it near by. Make sure there is enough service loop in the air line to remove the brake and set it aside.

- While holding the track brake housing, remove the track brake housing mounting bolts.  Do not remove the brake mounting plate.

# NOTE: There may be shims used to locate the brake.  Do not let the shims fall.  Note their location when removing the brake housing.

# Release the track brake calipers by applying air pressure to the brake.  Lift the track brake assembly off the brake housing mount and set in the manlift.  Insert the spacer between the brake pads and release the air pressure.

# NOTE: WATCH YOUR FINGERS.

- Remove the center portion of the flexible coupling.  Note the orientation and location of the spherical washers for assembly later.  Do not loose or drop the bolts and washers.  Keep the disc packs intact.  The roller screw shaft may rotate slightly due to the back driving force from the carriage.  Alternately, the track motor and coupling may be removed together off the roller screw.

- Loosen the screws that clamp the coupling to the roller screw shaft.  Remove the coupling flange attached to the roller screw shaft.  The coupling attached to the track motor can remain on the motor shaft.

NOTE: THE COUPLING IS KEYED TO THE SHAFT.

Do not drop the key.

- Loosen clamp screws on the track brake and slide the disc off the roller screw shaft.

NOTE: THE BRAKE DISC IS KEYED TO THE SHAFT.

Do not drop the key.

- Remove the fixed roller screws (+Y) bearing cover and v-seal from the fixed bearing mount.

- One person now is required at the floating bearing end (-Y) of the roller screw.

- At the floating bearing end (-Y), remove the bellows mounting screws and push the bellows in the +Y direction approximately 18 inches.

CAUTION: DO NOT DAMAGE THE THREADS ON THE ROLLER SCREW.

- At the floating bearing end (-Y), tightly wrap the roller screw with Mylar on the exposed threaded portion of the shaft.

- At the floating bearing end (-Y), locate the wooden v-block under the roller screw on the +Y side of the floating bearing mount (at the -Y end of the roller screw).

- Remove the v-seal and bearing cover from the floating bearing mount.

- At the fixed bearing end (+Y), attach the bearing retainer removal tool  (spanner) to the bearing retainer on the fixed bearing (+Y end).  Attach the ratchet driver.

- At the floating bearing end (-Y), attach the adjustable wrench on the flats on the end of the roller screw shaft.  Slip cheater bar over wrench handle for additional leverage.

WARNING: THE NEXT STEPS WILL RELEASE THE CARRIAGE AND ROLLER SCREW FROM THE FIXED SUPPORT.  MAKE SURE THE STRAP TO THE CRANE HOOK IS SECURED.

- Remove the Ringfeder clamp on the roller screw shaft.

- With one person securely holding the adjustable wrench, loosen the fixed bearing retainer by rotating the ratchet counter clockwise.  Remove the retainer.  Do not attempt to remove the bearings from the bearing mount.

- Do not attempt to remove the bearings or retainer from the roller screw on the floating roller screw end  (-Y end).

## 6.4 Y Roller-Screw Captive Bearing Removal

- The person at the lower end of the Y Roller-Screw (-Y) should rotate the roller-screw so that the screw moves upwards.  This will cause the fixed bearings to push out while the roller-screw moves through the floating bearings.  Insure that the roller-screw does not come free of the floating bearings or that floating bearings fall out of the support.

- Remove the fixed bearings.  Note the orientation and relationship.  Verify that these are correct.

- Alternative 1. Push the fixed bearings out of the fixed housing from the –Y side of the fixed housing.

6.5 Y Roller-Screw Captive Bearing Replacement

- Place the new bearings on the fixed end of the roller-screw in the proper orientation and relationship.

- Rotate the roller-screw from the lower end so that the fixed bearings seat in the housing and so that the roller-screw returns to the same position prior to step 6.5.  Insure that the floating bearings do not come free while the roller-screw is moving.

- Alternative 1. Locate the fixed bearings on the roller-screw and manually seat in the fixed housing.

- Insure that the fixed bearings are seated correctly in the fixed housing.

## 6.6 Y Roller-Screw Retainer Nut Replacement

Refer to P. 29-30 of the Tracker Installation Procedure (TM-12725). The remaining steps are an abridged version of that section of the procedure.  Also refer to drawing 1010-0020.

- Verify the shaft and bearings are correctly seated.  The floating bearings should be approximately centered in their housing.

- Perform steps in Note 12 on drawing 1010-0020.  Slide the bearing clamp on the roller screw shaft and snug up against the bearing.

- Attach the bearing retainer removal tool  (spanner) to the bearing retainer on the fixed bearing (+Y end).  Attach the torque wrench to the installation tool.

- At the floating bearing end (-Y), attach the adjustable wrench on the flats on the end of the roller screw shaft.  Slip the cheater bar over wrench handle for additional leverage.

- With one person securely holding the adjustable wrench, incrementally tighten the fixed bearing retainer by rotating the ratchet clockwise.  Use steps of 25, 50, 75, 100, 116 Ft-lbs.

- Install the Ringfeder clamp.

- Remove the wrenches and remaining Mylar.

- Slide the Cover, Retaining Nut on the drive end of the roller screw and seat on the fixed bearing outer race.  Install the screws with the flat and lock washers.  Torque to the required values (75 in-lbs) on 1010-0020. Use a “star pattern” when tightening.

- Apply a thin coat of silicone grease (DC-55 or equivalent) to the v-ring seal.  Slide the seal onto the roller screw shaft and seat against the bearing cover.

- Slide the Cover, Retaining Nut on the floating end of the roller screw shaft.  The cover should not touch the bearings.  Install the screws with flat and lock washers.  Torque to 40 in-lbs.

- Apply a thin coat of silicone grease (DC-55 or equivalent) to the v-ring seal.  Slide the seal onto the roller screw shaft and seat against the bearing cover.

The remaining assembly is the same as Sections 5.10, 5.11, and 5.12 in the Tracker Installation Procedure.  Except as noted below:

- Install the brake disc per Section 5.10

- Install the track motor if it was removed.  Check alignment and install the coupling per Section 5.11 steps b through i.

NOTE: Do not move the track motor bracket.  It should still be aligned.

- Install the track brake assembly using Section 5.12 steps b through f.

- Attach the roller screw bellows to their respective mounts.

- Attach all air lines that were removed.

The Y Roller-Screw Retainer Nut replacement is complete.

The track axis should be run to verify all components were reinstalled correctly and are aligned as follows.  In particular, listen for any clicking in the flexible coupling at the track motor / roller screw interface.  Clicking means the coupling is not aligned correctly.

- Verify all brakes are engaged.

- Reduce tension on the crane strap to allow for a 20 millimeter move in –Y.

- Test for alignment by commanding a –10 mm move in Y.  Listen for clicking from the track drive coupling.  If clicking is detected, notify the Installation Director.  Align the Y drive coupling per the alignment procedure and repeat this step.  Continue the alignment procedure and test until no clicking is heard from the Y track drive coupling.

- Disconnect the crane and safety straps from the carriage.